- Load Be+ (around 100s)

  - Parameters:

    - laser power(preferably smallest as possible throughout experiment)[1mW , 60mW] default 40mW (controlled via ramam board)

    - laser detuning [0MHz, -500MHz] default -300MHz

    - EC1 [5V, 50V] (default 10V)

    - EC2 [5V, 50V] (default 10V)

    - Comp_hor [0V, 50V] default 6V

    - Comp_ver [0V, 50V] default 37V

    - U_rf [0mV, 1500mV] default 700mV

    - B-field [on, off] (has to be on to load)

    - Be+ oven [on, off] (has to be on to load)

    - Bephi shutter [on, off](has to be on to load)

  - Sequence:

    - Set parameters optimal for Be+ loading

    - Preheat: Be+ oven[on]

    - Delay max 20s

    - Ionisation: Bephi shutter [on]

    - Delay max 60s

    - Be+ oven[off], Bephi shutter [off]

    - Set parameters for optimal cooling

  - Outcome:

    - Number of ions loaded

      - CCD image analysis

      - PMT signal analysis

  - comments

    - The system can stop the loading process anytime by stopping the Be+ oven and Bephi shutter

- Load HD+ (around 25s)

  - Parameters:

    - Piezo effect leak valve voltage(for controlling the HD concentration in the vacuum chamber) [0, 4V] default 2.4V

      - There is a set limit of pressure in the vacuum chamber, if it exceeds a certain limit, the valve has to be closed.

      - The smaller this value, the better for the experiment. Thus it is desired for the algorithm to minimize this value while optimizing loading success rate

    - Electron gun [on, off] (has to be on to ionise HD)

    - B-field [on, off] (has to be off to ionise HD, but has to be turned on again after loading HD)

    - UV3 laser power[1mW , 60mW]

    - UV3 laser detuning [0MHz, -500MHz] default -100MHz

    - EC1 [5V, 50V] default 30V

    - EC2 [5V, 50V] default 30V

    - Comp_hor [0V, 50V] default 8V

    - Comp_ver [0V, 50V] default 43V

    - U_rf [0mV, 1500mV] default 100V

  - Sequence: (!!!!!!!!if anything calls an error, cloes the HD valve and egun immediately!!!!!)

    - Set parameters for optimal HD+ loading

    - Preheat e-gun: E-gun [on]

    - Delay max (10s)

    - Load HD:

      - Set piezo voltage [0, 10V]

      - Open valve [on]

    - Delay max 5s

    - Ionisation: B-field[off]

    - Delay max 5s

    - Stop: egun[off], valve [off], B-field[on]

    - Set parameters for optimal cooling

  - Outcome(if able to detect):

    - If 0 dark ion loaded (fail, repeat step 2)

    - If 1 dark ion loaded (success, proceed to step 3)

    - If more than 1 ion loaded (fail, set U_rf[0mV], then repeat step 1)

- Ejection (from 0s to 60s)

  - Parameters:

    - laser power[1mW , 60mW] default 1mW

    - laser detuning [0MHz, -500MHz] default -200MHz

    - EC1 [5V, 50V] default 30V

    - EC2 [5V, 50V] default 30V

    - Comp_hor [0V, 50V] default 8V

    - Comp_ver [0V, 50V] default 43V

    - U_rf [0mV, 1500mV] default 400mV

  - Sequence:

    - Count Be+ ions

      - If 0 Be+ ion, set U_rf[0mV], then go back step 1

      - If 1 Be+ ion, proceed to step 4

      - If more than 1 Be+ ion: proceed to step 3.b.ii.

    - Set parameter for the specific Be+ number for ejection

    - Delay max 30s

    - Set parameter for optimal cooling

    - Repeat step 3.b.i.

- Test secular frequency

  - Set parameter for optimal cooling and micromotion minimization

  - Calculate secular frequencies for current parameters

  - Axial scan

    - If scan result = 4.b. value (within accepted uncertainty): SUCCESS

    - If scan result != 4.b. value: FAIL (set u_rf [0mV] and restart from step 1)

Here is the formal technical documentation for your MHI loading optimization system. This document focuses on the algorithmic logic, control strategy, and theoretical framework, omitting specific code implementations.

Automated Optimization of Multi-Species Ion Loading via Trust Region Bayesian Optimization (TuRBO)

1. Introduction & Objective

The objective of this system is to automate the parameter optimization for the loading of a mixed Coulomb crystal consisting of one Beryllium ion ($^{9}\text{Be}^+$) and one Hydrogen Deuteride ion ($\text{HD}^+$).

The physical system presents a high-dimensional control problem ($14$ continuous parameters) characterized by:

- Stochasticity: Loading is a probabilistic Poisson process.

- High Latency: A single experimental cycle (load-clean-verify) can require up to 10 minutes.

- Coupled Physics: The efficiency of loading the target ion ($\text{HD}^+$) is non-linearly dependent on the initial number of coolant ions ($^{9}\text{Be}^+$) loaded in the previous stage.

This document outlines the implementation of an "Experiment-in-the-Loop" optimization architecture that replaces manual tuning with a sample-efficient Bayesian optimizer.

2. Algorithmic Framework

2.1 Core Algorithm: TuRBO

We utilize TuRBO (Trust Region Bayesian Optimization) rather than standard Reinforcement Learning (RL) or Grid Search.

- Surrogate Modeling (Gaussian Process): Unlike RL, which requires millions of trials to learn a policy, TuRBO builds a probabilistic model of the parameter landscape. It predicts the "Loading Quality" for untested parameters based on previous results.

- Trust Regions: To manage the high-dimensionality (30+ params), the algorithm restricts optimization to a local "Trust Region" (a dynamic hyper-rectangle).

  - Success: If a trial improves the result, the region expands (exploration).

  - Failure: If a trial fails, the region shrinks (exploitation).

  - Restart: If the region becomes too small without success, the algorithm resets to a new area of the parameter space to avoid local minima.

2.2 Safety Constraints

The algorithm operates within strict "Hard Bounds" to prevent hardware damage:

- Oven Current: Clamped to maximum safe operational limits.

- Piezo Leak Valve: Constrained to $0V - 10V$.

- RF Power: Limited to prevent breakdown voltages.

3. Optimization Strategy: The Hierarchical Approach

A naive optimization of all 30 parameters simultaneously is infeasible due to the "Thermal Mass vs. Cleaning" trade-off:

More $\text{Be}^+$ ions facilitate $\text{HD}^+$ capture (sympathetic cooling) but impede the final purification (ejection) process.

To solve this, we employ a Two-Phase Hierarchical Strategy:

Phase I: Ejection Stability (Pre-requisite Calibration)

Before optimizing loading, the system must be robust against "over-loading."

- Logic: Intentionally load excess $\text{Be}^+$ ions (5-10) and optimize only the Step 3 (Ejection) parameters.

- Goal: Find static parameters that reliably reduce $N \approx 5$ ions to $N = 1$ ion.

- Result: These parameters are locked during Phase II.

Phase II: End-to-End Loading Optimization (Main Loop)

The algorithm optimizes Step 1 (Be+ Load) and Step 2 (HD+ Load) simultaneously as a single coupled chain.

- Dynamic Cooling Logic: The parameters for Step 2 are not static. The cooling laser power ($P_{cool}$) is parameterized as a function of the detected $\text{Be}^+$ count ($N_{Be}$):

$$P_{cool} = P_{base} + (k \cdot N_{Be})$$

The optimizer tunes $P_{base}$ and the scaling factor $k$, allowing it to learn different cooling strategies for different crystal sizes.

4. The Objective Function (Reward Logic)

The optimizer maximizes a scalar Reward Score ($R$) calculated at the end of every cycle. The score balances three competing objectives: Success, Vacuum Quality, and Speed.

The formula is defined as:

$$R = R_{Outcome} - C_{Pressure} + R_{Speed}$$

A. Outcome Reward ($R_{Outcome}$)

Determined by the final ion configuration after Step 3.

- Target State ($1\text{Be}^+ + 1\text{HD}^+$): +200 (Global Maxima).

- Partial Success ($1\text{Be}^+ + 0\text{HD}^+$): -10 (Failure to load, but crystal preserved).

- Catastrophic Failure (0 ions or $>1$ Dark ions): -100 (Requires full trap dump).

B. Pressure Cost ($C_{Pressure}$)

This term forces the agent to minimize the Piezo voltage ($V_{piezo}$) to maintain UHV conditions.

$$C_{Pressure} = \alpha \cdot V_{piezo}$$

(Where $\alpha$ is a weighting factor determined by vacuum sensitivity).

C. Speed Bonus ($R_{Speed}$)

To optimize the "Time-to-Load" within the 10-minute window.

$$R_{Speed} = \frac{T_{Max} - T_{Elapsed}}{T_{Max}} \times 50$$

- Effect: A fast load (1 min) yields a significantly higher score than a slow load (9 mins), even if both result in success.

5. Process Control Logic

The following logic governs the execution of a single experimental trial.

Step 1: Initialization & Be+ Loading

- Laser Lock Check: If lasers are unlocked, pause optimization and run re-lock sequence.

- Apply Parameters: Set Oven, RF, and Lasers based on TuRBO's suggestion.

- Load: Run Be+ loading sequence.

- Check:

  - If $N_{Be} == 0$: Abort. Return Reward $-10$ (Wait time penalty).

  - If $N_{Be} > 0$: Proceed (Do not dump excess ions).

Step 2: HD+ Sympathetic Loading

- Calculate Cooling: Adjust cooling parameters based on $N_{Be}$.

- Stopwatch Start: Begin 10-minute maximum window.

- Loop:

  - Open Piezo Valve / Pulse E-Gun.

  - Monitor Fluorescence (PMT/Camera).

  - Early Stop Condition: If a "dark event" (drop in fluorescence or mass change) is detected, stop immediately.

- Timeout: If $T > 10\text{ mins}$, stop and proceed to cleanup.

Step 3: Purification (Ejection)

- Apply Fixed Phase I Parameters: Use the robust ejection settings derived in Phase I.

- Sequence: Attempt to eject excess $\text{Be}^+$ and impurity ions, leaving only the target $1\text{Be}^+ + 1\text{HD}^+$ pair.

Step 4: Verification & Scoring

- Secular Frequency Scan: Verification of the dark ion mass.

- Classification: Assign the trial to one of three categories:

  - Success (Target reached).

  - Miss (No dark ion).

  - Fail (Wrong mass / Trap empty).

- Feedback: Compute $R$ and send to TuRBO to update the Gaussian Process model.

6. Visual Logic Flowchart

Code snippet

flowchart TD

%% Nodes

Start([Start Optimization Loop])

Turbo[<b>TuRBO Algorithm</b><br/>Generates Parameters]

Safety{Safety Check<br/>Within Bounds?}

Lock{Lasers Locked?}

Step1[<b>Step 1: Be+ Load</b><br/>Variable N count]

CheckBe{N_Be > 0?}

Step2[<b>Step 2: HD+ Load</b><br/>Dynamic Cooling: P(N_Be)<br/>Stopwatch Mode]

Stop{Ion Detected OR<br/>Timeout?}

Step3[<b>Step 3: Ejection</b><br/>Uses Fixed Robust Params]

Verify[<b>Step 4: Verification</b><br/>Secular Scan]

CalcReward[<b>Calculate Reward</b><br/>R = Outcome - Pressure + Speed]

Update[Update Gaussian Model]

%% Connections

Start --> Turbo

Turbo --> Safety

Safety -- No --> Turbo

Safety -- Yes --> Lock

Lock -- No --> Relock[Relock Routine] --> Lock

Lock -- Yes --> Step1

Step1 --> CheckBe

CheckBe -- No --> Fail1[Reward: -10] --> Update

CheckBe -- Yes --> Step2

Step2 --> Stop

Stop -- Yes --> Step3

Step3 --> Verify

Verify --> CalcReward

CalcReward --> Update

Update --> Turbo

Blind spots in the experiment for manual calibration

- UV3 laser lock optimisation

- Bephi optimisation

- E-gun setting optimisation

feed back parameters for ion counting:

- N

- x_i

- y_i

- sig_x_i

- sig_y_i

- area_i

- intensity_i